## Requirements
Controller: |  | PCI | [5 empty] | Controller | PSIM | Cables
[8 empty] | PCI | AC PSIM | Orbit Cable (006869)
Supply I Max | Controller Requires PSIM | High Speed Capable | Dynamic Capable | Notes |  |  |  | USB | DC PSIM | Orbit Cable (804731-xxx)
1000 | No | Yes | Yes | Though the PCI Card can supply up to 31 probes that is only if the computer is build to the pci specifications correctly with regards power therefore we recommend people only draw 1000mA from the card to ensure it works correctly |  |  |  | RS232 | 5V PSIM
Preferred PSIM if Required: |  | AC PSIM
Imax | Output Impedance | Output V |  | Yellow cells can be changed
1600 | 145.2 | 5.2
Modules Required
Module |  | Num Required | Typical Load Current for 1 Module |  | Total Current Required
DIOM |  | 0 | 53.7 |  | 0 |  |  | Start V
EIM |  | 0 | 79.3 |  | 0 |  |  | 5.1
DP (Includes all displacement modules) |  | 0 | 59 |  | 0
DIM |  | 0 | 40.3 |  | 0 |  |  | Max I inc Vdrop
AIM - PT100 |  | 0 | 73.6 |  | 0 |  |  | 2187.5 | 1000
AIM 4-20 mA |  | 0 | 154 |  | 0 |  |  | 1000
AIM Voltage |  | 0 | 77.8 |  | 0
Happy Light |  | 0 | 8.1 |  | 0
LE-EIM |  | 0 | 134.4 |  | 0 |  |  | Required Num PSIMS
LE |  | 0 | 41.6 |  | 0 |  |  | 0 | 0
Total: |  | 0 | - |  | 0 |  |  | 0
Cable Type: |  | Orbit Cable (804731-xxx) |  |  | Max Current from controller allowing a max Volt Drop to 4.75V: | 1000 | mA
Required Length: |  | 10 |  | Meters
Total Cable Resistance: |  | 0.16 |  | Ohms | Required Number of PSIMS: | 0
(This assumes ambient temperature not greater than 40degC, See PSIM section of the orbit system manual for derating)

## Sheet1 (hidden)
Controller Source | Imax | Requires PSIM | Notes | High Speed Capable | Dynamic Capable
PCI | 1000 | No | Though the PCI Card can supply up to 31 probes that is only if the computer is build to the pci specifications correctly with regards power therefore we recommend people only draw 1000mA from the card to ensure it works correctly | Yes | Yes | 5.1
USB | 200 | No |  | No | No | 5.05
RS232 | -65.5 | Yes | This controller requires a PSIM as power is not supplied down RS232 | No | No | 0
Unit | V |  | I Min Idle | I min Reading | I Typical Idle (Loaded) | I Typical Reading (Loaded) | I Max Idle | I Max Reading | Note
DIOM | 5.1 |  | 28.3 | 41.3 | 40.2 | 53.7 | 428.3 | 441.3 | To work out the Exact DIOM power requirements take the minimim reading value + 8 * the maximum source/sync current you require on a pin
EIM | 5.1 |  | 34.2 | 49 | 64.6 | 79.3 | 1034.2 | 1049 | The Exact power requirement is the Minimum Read Current + the encoder drive current.
DP (Includes all displacement modules) | 5.1 |  | 46 | 59 | 46 | 59 | 46 | 59
DIM | 5.1 |  | 25.8 | 39.5 | 25.9 | 40.3 | 25.9 | 40.3 | May vary slightly with DIM devices Connected
AIM - PT100 | 5.1 |  | 69.1 | 73.6 | 69.1 | 73.6 | 69.1 | 73.6
AIM 4-20 mA | 5.1 |  | 70.3 | 75.2 | 142.7 | 154 | 142.7 | 154
AIM Voltage | 5.1 |  | 68.5 | 77.8 | 68.5 | 77.8 | 68.5 | 77.8
Happy Light | 5.1 |  | 8.1 | 8.1 | 8.1 | 8.1 | 8.1 | 8.1
LE-EIM | 5.1 |  | 134.4 | 134.4 | 134.4 | 134.4 | 134.4 | 134.4
LE | 5.1 |  | 40 | 41.6 | 40 | 41.6 | 40 | 41.6
PSIM | Imax | Output Impedance | Vout
AC PSIM | 1600 | 145.2 | 5.2
DC PSIM | 1600 | 66.1 | 5.14
5V PSIM | 1000 | 327 | 5
CableType | R/M
Orbit Cable (006869) | 0.1
Orbit Cable (804731-xxx) | 0.016
Sum Of Current Reqirement | Vsupply |  |  | Min V | Effective R of modules | Cable Rpm | Cable Length | Rcable | Max Load resistance at end of cable | Max Current Drawable from Supply (A)
100 | 5.1 |  |  | 4.75 | 0.0475 | 0.016 | 10 | 0.16 | 2.1714 | 2.1875